STEREO

IMPACT

HET Comprehensive Performance Test Procedure

HET_CPT_D.DOC

Rev. D,  06/08/05

Larry Ryan

Tycho von Rosenvinge

Kristin Wortman

Date Run: ______________________________

Flight Model Location: ___________________

Flight Model Tested: _____________________

Flight Software Version:__________________

Data File Recorded:______________________

Performed By: __________________________

- Test Description

This comprehensive test procedure has been designed to run on either HET flight model (FM1 and FM2).  The WinMac-GSE software is required to send the appropriate commands for configuration and to view the HET packets to extract the relevant information for verification.

HET should be powered with the high voltage bias applied to the detectors, and the HET flight code should be booted with the HET tables successfully loaded.  The HET flight model should be in default boot configuration before running any test.  If previous tests were conducted, then HET should be rebooted before running any further tests.

The results of these tests will verify that the HET flight model is operational and the flight software is accepting and processing commands.

- Housekeeping Data

Record the housekeeping data for the seven HET detectors.  If HET has just been booted, it will take 5-7 minutes for all housekeeping data to come through.

Start Time: _____   HET Temperature #1 _____   HET Temperature #2 _____

Data from HK (dec) page:

|  | H1i | H1o | H2 | H3 | H4 | H5 | H6 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| PHASIC |  |  |  |  |  |  |  |
| Channel |  |  |  |  |  |  |  |
| Preamp |  |  |  |  |  |  |  |
| HG Thresh |  |  |  |  |  |  |  |
| LG Thresh |  |  |  |  |  |  |  |
| Leakage |  |  |  |  |  |  |  |
| DAC |  |  |  |  |  |  |  |

The following are the expected values for FM1 and FM2:

|  | H1i | H1o | H2 | H3 | H4 | H5 | H6 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| PHASIC | 0 | 0 | 1 | 1 | 1 | 1 | 1 |
| Channel (FM1) | 1 | 11 | 3 | 5 | 7 | 9 | 12 |
| Channel (FM2) | 1 | 10 | 2 | 5 | 7 | 10 | 13 |
| Preamp (volts) | ~4.40 | ~4.40 | ~4.40 | ~4.40 | ~4.40 | ~4.40 | ~4.40 |
| HG Thresh (FM1) | 284 | 220 | 36 | 188 | 140 | 180 | 100 |
| HG Thresh (FM2) | 192 | 244 | 124 | 120 | 68 | 68 | 30 |
| LG Thresh (FM1) | 176 | 192 | 112 | 8 | 96 | 112 | 248 |
| LG Thresh (FM2) | 192 | 192 | 44 | 56 | 64 | 52 | 52 |
| Leakage (uA) | <0.5 | <0.5 | <0.5 | <2.0 | <2.0 | <2.0 | <1.0 |

3.  Singles Rates / Spurious Events

Record the H1-H6 singles rates (low gain and high gain) for three consecutive major frames.  The high gain singles rates should be < 100/min., and the low gain singles rates should be < 5/min.

|  | High-Gain Singles Rates for 3 major frames | Low-Gain Singles Rates for 3 major frames |
| --- | --- | --- |
| Mjr Fr # |  |  |  |  |  |  |
| H1i |  |  |  |  |  |  |
| H1o |  |  |  |  |  |  |
| H2 |  |  |  |  |  |  |
| H3 |  |  |  |  |  |  |
| H4 |  |  |  |  |  |  |
| H5 |  |  |  |  |  |  |
| H6 |  |  |  |  |  |  |

Record the following rates for one major frame.    Major Frame #: ____________

| Rate Counter | Expected Counts | Actual Counts | Rate Counter | Expected Counts | Actual Counts |
| --- | --- | --- | --- | --- | --- |
| Livetime | ~12 x 106 |  | Stop heavies | 0 |  |
| Trigger | <300 |  | Pen H | 0 |  |
| Coincidence | 0 |  | Pen He | 0 |  |
| Events | 0 |  | Pen heavies | 0 |  |
| Singles que | 0 |  | Invalid | 0 |  |
| Stop que | 0 |  | H1 I and O | 0 |  |
| Pen que | 0 |  | Inconsist. | 0 |  |
| Stop H | 0 |  | H1 not first | 0 |  |
| Stop He | 0 |  | Stim evt cnt | 0 |  |

4.  Event Identification / Commanding

4.1.  Stopping H Events

Send the following HET commands:

FM1					FM2

tpfbsel 0 b 116			tpfbsel 0 a 116

tpfbsel 1 3 116			tpfbsel 1 2 116

testp 4e2 2d 1e			testp 4e2 2d 1e

Record the following rates counters.  All other rate counters should read 0.

Major Frame #  ___________________

| Rate Counter | Expected Counts | Actual Counts |
| --- | --- | --- |
| Coincidence | ~600 |  |
| Stop que | ~600 |  |
| Stop H | ~600 |  |
| Stim evt cnt | ~600 |  |

4.2.  Stopping He Events

Send the following HET commands:

FM1					FM2

tpfbsel 0 b 6			tpfbsel 0 a 6

tpfbsel 0 1 116			tpfbsel 0 1 116

tpfbsel 1 5 12c			tpfbsel 1 5 12c

testp 4e2 10f 1e			testp 4e2 10f 1e

Record the following rates counters.  All other rate counters should read 0.

Major Frame #  ___________________

| Rate Counter | Expected Counts | Actual Counts |
| --- | --- | --- |
| Coincidence | ~600 |  |
| Stop que | ~600 |  |
| Stop He | ~600 |  |
| Stim evt cnt | ~600 |  |

4.3.  Stopping Z>2 Events

Send the following HET commands:

FM1					FM2

tpfbsel 0 1 6				tpfbsel 0 1 6

tpfbsel 0 b 116			tpfbsel 0 a 116

tpfbsel 1 5 11c			tpfbsel 1 5 11c

testp 4e2 179 3d			testp 4e2 179 3d

Record the following rates counters.  All other rate counters should read 0.

Major Frame #  ___________________

| Rate Counter | Expected Counts | Actual Counts |
| --- | --- | --- |
| Coincidence | ~600 |  |
| Stop que | ~600 |  |
| Stop heavies | ~600 |  |
| Stim evt cnt | ~600 |  |

4.4.  Penetrating H Events

Send the following HET commands:

FM1					FM2

tpfbsel 1 3 126			tpfbsel 1 2 126

tpfbsel 1 7 11c			tpfbsel 1 7 11c

tpfbsel 1 9 11c			tpfbsel 1 a 11c

tpfbsel 1 c 116			tpfbsel 1 d 116

testp 4e2 17 b			testp 4e2 16 a

Record the following rates counters.  All other rate counters should read 0.

Major Frame #  ___________________

| Rate Counter | Expected Counts | Actual Counts |
| --- | --- | --- |
| Coincidence | ~600 |  |
| Pen que | ~600 |  |
| Pen H | ~600 |  |
| Stim evt cnt | ~600 |  |

- Stimulus Pulser / Analog Channels  (FM1 only, skip to step 6 for FM2)

First, send the following HET commands:

gwrite 9 32

tmode 3

hgorsel 1 0 12a8

lgorsel 1 0 12a8

The next section has been designed to pulse all the channels (H1i, H1o, H2, H3, H4, H5, H6).    The command set should be sent and the data values should be reviewed and recorded at each configuration as directed.  It should be noted that the expected pulse height value range may not be exact, but the values should lie very close to this range.

5.1.  Configure All Channels

Send the following HET command set to configure the channels as follows:

- enable the test pulser for all channels

- set the feedback capacitor to 30pF  for H1i, H1o, H2 and H6

- set the feedback capacitor to 60pF  for H3, H4 and H5

- set the test input capacitor to 18pF for H1i, H1o, H2 and H6

- set the test input capacitor to 36pF for H3, H4 and H5

tpfbsel 0 1 166

tpfbsel 0 b 166

tpfbsel 1 3 166

tpfbsel 1 5 1cc

tpfbsel 1 7 1cc

tpfbsel 1 9 1cc

tpfbsel 1 c 166

Send the following HET command set to configure the channels as follows:

- disable the high gain ADC for all channels

- enable the low gain ADC for all channels

phacont 0 1 7

phacont 0 b 7

phacont 1 3 7

phacont 1 5 7

phacont 1 7 7

phacont 1 9 7

phacont 1 c 7

Send the following HET command to configure the following:

- set the test pulser frequency to 10 pulses per second

- set the Phasic 0 DAC setting to 5V

- set the Phasic 1 DAC setting to 5V

testp 4e2 1ff 1ff

Wait about 2-3 minutes and then record the range of raw data observed (a 6-digit hex number, the last three of which represent the actual pulse height) using raw data (pulse heights from packet 597:

| Channel | Expected Range | Actual Range | Mjr Fr #s |
| --- | --- | --- | --- |
| H1i | 08c813 – 08c818 |  |  |
| H1o | 05c80b – 05c80f |  |  |
| H2 | 11c804 – 11c808 |  |  |
| H3 | 12c830 – 12c834 |  |  |
| H4 | 13c818 – 13c81d |  |  |
| H5 | 14c824 – 14c828 |  |  |
| H6 | 96481c – 964821 |  |  |

Send the following HET command to configure the following:

- set the test pulser frequency to 10 pulses per second

- set the Phasic 0 DAC setting to 2.51V

- set the Phasic 1 DAC setting to 2.51V

testp 4e2 180 180

Wait about 2-3 minutes and then record the range of raw data observed (a 6-digit hex number, the last three of which represent the actual pulse height) using raw data (pulse heights from packet 597:

| Channel | Expected Range | Actual Range | Mjr Fr #s |
| --- | --- | --- | --- |
| H1i | 08c423 – 08c425 |  |  |
| H1o | 05c41d – 05c420 |  |  |
| H2 | 11c41a – 11c41c |  |  |
| H3 | 12c42e – 12c430 |  |  |
| H4 | 13c422 – 13c425 |  |  |
| H5 | 14c428 – 14c42a |  |  |
| H6 | 964425 – 964427 |  |  |

Send the following command set to configure all the channels to the following:

- enable the high gain ADC for all channels

- enable the low gain ADC for all channels

- verbose mode off

phacont 0 1 d

phacont 0 b d

phacont 1 3 d

phacont 1 5 d

phacont 1 7 d

phacont 1 9 d

phacont 1 c d

Send the following HET command to configure the following:

- set the test pulser frequency to 10 pulses per second

- set the Phasic 0 DAC setting to .25V

- set the Phasic 1 DAC setting to .25V

testp 4e2 ff ff

Wait about 2-3 minutes and then record the range of raw data observed (a 6-digit hex number, the last three of which represent the actual pulse height) using raw data (pulse heights from packet 597:

| Channel | Expected Range | Actual Range | Mjr Fr #s |
| --- | --- | --- | --- |
| H1i | 0887e7 – 0887eb |  |  |
| H1o | 0587cb – 0587d0 |  |  |
| H2 | 1187eb – 1187f0 |  |  |
| H3 | 1287ec – 1287f0 |  |  |
| H4 | 1387e8 – 1387ed |  |  |
| H5 | 1487de – 1487e2 |  |  |
| H6 | 9607f6 – 9607fb |  |  |

Send the following HET command to configure the following:

- set the test pulser frequency to 10 pulses per second

- set the Phasic 0 DAC setting to 0.125V

- set the Phasic 1 DAC setting to 0.125V

testp  4e2 80 80

Wait about 2-3 minutes and then record the range of raw data observed (a 6-digit hex number, the last three of which represent the actual pulse height) using raw data (pulse heights from packet 597:

| Channel | Expected Range | Actual Range | Mjr Fr #s |
| --- | --- | --- | --- |
| H1i | 08840b – 088410 |  |  |
| H1o | 0583fd – 058401 |  |  |
| H2 | 11840c – 118410 |  |  |
| H3 | 12840c – 128410 |  |  |
| H4 | 13840a – 13840f |  |  |
| H5 | 148405 – 148409 |  |  |
| H6 | 960411 – 960416 |  |  |

Send the following HET command to configure the following:

- set the test pulser frequency to 10 pulses per second

- set the Phasic 0 DAC setting to .002V

- set the Phasic 1 DAC setting to .002V

testp 4e2 2 2

Wait about 2-3 minutes and then record the range of raw data observed (a 6-digit hex number, the last three of which represent the actual pulse height) using raw data (pulse heights from packet 597:

| Channel | Expected Range | Actual Range | Mjr Fr #s |
| --- | --- | --- | --- |
| H1i | 08803d - 088040 |  |  |
| H1o | 05803b – 05803f |  |  |
| H2 | 11803b – 11803e |  |  |
| H3 | 12803a – 12803e |  |  |
| H4 | 13803b – 13803e |  |  |
| H5 | 14803a – 14803e |  |  |
| H6 | 96003a – 96003e |  |  |

Finally, send the following HET commands or reboot:

testp 4e2 0 0

gwrite 9 30

tmode 0

hgorsel 1 0 8

lgorsel 1 0 8

tpfbsel 0 1 6

tpfbsel 0 b 6

tpfbsel 1 3 6

tpfbsel 1 5 c

tpfbsel 1 7 c

tpfbsel 1 9 c

tpfbsel 1 c 6

*** END OF TEST FOR FM1.  STOP HERE IF TESTING FM1.***

6.  Stimulus Pulser / Analog Channels  (FM2 only)

First, send the following HET commands:

gwrite 9 32

tmode 3

hgorsel 1 0 24a4

lgorsel 1 0 24a4

6.1.   Configure All Channels

Send the following HET command set to configure the channels as follows:

- enable the test pulser for all channels

- set the feedback capacitor to 30pF  for H1i, H1o, H2 and H6

- set the feedback capacitor to 60pF  for H3, H4 and H5

- set the test input capacitor to 18pF for H1i, H1o, H2 and H6

- set the test input capacitor to 36pF for H3, H4 and H5

tpfbsel 0 1 166

tpfbsel 0 a 166

tpfbsel 1 2 166

tpfbsel 1 5 1cc

tpfbsel 1 7 1cc

tpfbsel 1 a 1cc

tpfbsel 1 d 166

Send the following HET command set to configure the channels as follows:

- disable the high gain ADC for all channels

- enable the low gain ADC for all channels

phacont 0 1 7

phacont 0 a 7

phacont 1 2 7

phacont 1 5 7

phacont 1 7 7

phacont 1 a 7

phacont 1 d 7

Send the following HET command to configure the following:

- set the test pulser frequency to 10 pulses per second

- set the Phasic 0 DAC setting to 5V

- set the Phasic 1 DAC setting to 4.71V

testp 4e2 1ff 1f0

Wait about 2-3 minutes and then record the range of raw data observed (a 6-digit hex number, the last three  of which represent the actual pulse height)  using raw data (pulse heights from packet 597:

| Channel | Expected Range | Actual Range | Mjr Fr #s |
| --- | --- | --- | --- |
| H1i | 08c816 – 08c818 |  |  |
| H1o | 054808 – 05480d |  |  |
| H2 | 1147e8 – 1147ed |  |  |
| H3 | 12c7ea – 12c7ef |  |  |
| H4 | 13c7f0 – 13c7f5 |  |  |
| H5 | 1547ea – 1547ef |  |  |
| H6 | 96c7dc – 96c7e1 |  |  |

Send the following HET command to configure the following:

- set the test pulser frequency to 10 pulses per second

- set the Phasic 0 DAC setting to 2.51V

- set the Phasic 1 DAC setting to 2.51V

testp 4e2 180 180

Wait about 2-3 minutes and then record the range of raw data observed (a 6-digit hex number, the last three of which represent the actual pulse height) using raw data (pulse heights from packet 597:

| Channel | Expected Range | Actual Range | Mjr Fr #s |
| --- | --- | --- | --- |
| H1i | 08c423 – 08c426 |  |  |
| H1o | 05441c– 05441e |  |  |
| H2 | 11444b – 11444e |  |  |
| H3 | 12c44c – 12c44e |  |  |
| H4 | 13c44f – 13c451 |  |  |
| H5 | 15444c – 15444e |  |  |
| H6 | 96c444 – 96c447 |  |  |

Send the following command set to configure all the channels to the following:

- enable the high gain ADC for all channels

- enable the low gain ADC for all channels

- verbose mode off

phacont 0 1 d

phacont 0 a d

phacont 1 2 d

phacont 1 5 d

phacont 1 7 d

phacont 1 a d

phacont 1 d d

Send the following HET command to configure the following:

- set the test pulser frequency to 10 pulses per second

- set the Phasic 0 DAC setting to .25V

- set the Phasic 1 DAC setting to .235V

testp 4e2 ff f0

Wait about 2-3 minutes and then record the range of raw data observed (a 6-digit hex number, the last three of which represent the actual pulse height) using raw data (pulse heights from packet 597:

| Channel | Expected Range | Actual Range | Mjr Fr #s |
| --- | --- | --- | --- |
| H1i | 0887cc – 0887d1 |  |  |
| H1o | 0507d2 – 0507d6 |  |  |
| H2 | 1107c9 – 1107cd |  |  |
| H3 | 1287ad – 1287b1 |  |  |
| H4 | 1387af – 1387b4 |  |  |
| H5 | 1507b4 – 1507b8 |  |  |
| H6 | 9687a3 – 9687a8 |  |  |

Send the following HET command to configure the following:

- set the test pulser frequency to 10 pulses per second

- set the Phasic 0 DAC setting to .125V

- set the Phasic 1 DAC setting to .125V

testp  4e2 80 80

Wait about 2-3 minutes and then record the range of raw data observed (a 6-digit hex number, the last three of which represent the actual pulse height) using raw data (pulse heights from packet 597:

| Channel | Expected Range | Actual Range | Mjr Fr #s |
| --- | --- | --- | --- |
| H1i | 0883fd – 088402 |  |  |
| H1o | 0503ff – 050404 |  |  |
| H2 | 110439 – 11043e |  |  |
| H3 | 12842a – 12842f |  |  |
| H4 | 13842b – 138430 |  |  |
| H5 | 15042e – 150432 |  |  |
| H6 | 968425 – 968429 |  |  |

Send the following HET command to configure the following:

- set the test pulser frequency to 10 pulses per second

- set the Phasic 0 DAC setting to .002V

- set the Phasic 1 DAC setting to .002V

testp 4e2 2  2

Wait about 2-3 minutes and then record the range of raw data observed (a 6-digit hex number, the last three of which represent the actual pulse height) using raw data (pulse heights from packet 597:

| Channel | Expected Range | Actual Range | Mjr Fr #s |
| --- | --- | --- | --- |
| H1i | 08803c – 08803f |  |  |
| H1o | 05003b – 05003e |  |  |
| H2 | 11003d – 110041 |  |  |
| H3 | 12803d – 128041 |  |  |
| H4 | 13803d – 138040 |  |  |
| H5 | 15003d – 150040 |  |  |
| H6 | 96803c – 96803f |  |  |

Finally, send the following HET commands or reboot:

testp 4e2 0 0

gwrite 9 30

tmode 0

hgorsel 1 0 4

lgorsel 1 0 4

tpfbsel 0 1 6

tpfbsel 0 a 6

tpfbsel 1 2 6

tpfbsel 1 5 c

tpfbsel 1 7 c

tpfbsel 1 a c

tpfbsel 1 d 6