STEREO/IMPACT/SIT

Comprehensive Performance Test

Test Procedure

Ver. 3.2  05/12/05

(Revised by K. Wortman)

Unit:			Test:							Date:

Recorded by:  _______________________________    Time: ____

1.  PURPOSE

This document specifies the procedure for performing the SIT Comprehensive Performance Test (CPT) whenever needed.  The CPT is designed to provide as large as possible, but not total, test of the performance capabilities of the SIT sensor.  It is intended for use in all environments in which access to the instrument is possible.  Certain portions of the sensor, e.g. the HVPS, and certain aspects of performance, e.g. calibration, are not tested because this can only be done in limited environments: the HVPS can only be turned on in high vacuum and the calibration requires significant amounts of bench test equipment.  The CPT should be performed whenever the full operation of SIT needs to be verified, as opposed to simple liveness checks.

2.  REQUIREMENTS

2.1 Equipment – A battery-powered SIT Test Pulser is required for this test, to be connected to the test connectors.

3. PROCEDURE

NOTE: It is unnecessary to repeat commands between procedures if they would simply put SIT in the state it is already in.

3.1 Check Test Pulser and Install

NOTE:  This section may be performed independently of the state of SIT power.

a) Switch Test Pulser ON/OFF Switch to ON

b) Press TEST pushbutton.  The green BATTERY LED should light.  The red LOW VOLTAGE LED should not light.  Absence of green LED or presence of red LED indicates a worn out battery, so replace it.

c) Switch ON/OFF switch to OFF

d) Verify that the SIT telescope cover is closed.

e) Ground the Test Pulser chassis to the test setup and place the pulser in a stable location within the pulser’s 1 meter cable length of the SIT instrument.

f) Connect the three signal coax’s from the pulser to the SIT test inputs.  See Fig 1.

NOTE:  IN CASE Test Pulser not available – connect bench pulsers as in Fig 1A.  The DG535 should be powered off for the beginning of the test.

3.2 Instrument Power On and Initialization

- If loading code from SEP Central EEPROM, verify EEPROM checksum before proceeding.

| Description | Expected Value | Actual Value |
| --- | --- | --- |
| tab cksm | 927143 (04/25/05) |  |

- If loading code from a file, verify correct file being used.

c)  Power up SIT if necessary and boot.  Verify it has properly booted.

Begin login of SIT data into a dated file.

Log Filename  __________________________________

d) Verify SEP Bias supply is ON

e) Send the SIT Command

immed 1

f)  Verify data packets arriving in range 605-619, and that sequence counts are reasonable and that the major frame number advances.

- After at least 2 minutes, verify incoming data:

| Description | Expected Value | Actual Value | Comments |
| --- | --- | --- | --- |
| SRT | 0 |  |  |
| STP | 0 |  |  |
| VS | 0 |  |  |
| SSD | <20 |  |  |
| VSE | 0 |  |  |
| Matrix Rates (ApId 605) | All 0 |  |  |
| Beacon Rates (ApId 619) | All 0 |  |  |
| Pulse Height Events (ApId 606-617) | All 0 |  |  |

i)  After 3-5 minutes verify the following status information has been returned and matches expected values:

| Description | Expected Value | Actual Value | Comments |
| --- | --- | --- | --- |
| SW Ver | 0903 |  |  |
| SW error | 0 = (no error) |  |  |
| HV Step | 00 |  |  |
| Toferror | 1 = error events |  |  |
| HV Status | 0 = off |  |  |
| Eonly | 0 = ET coincidence required |  |  |
| Junk (Events) | 0 = junk events ignored |  |  |
| LIMHI | 500 |  |  |
| calib gain | 9 - 11 |  |  |
| calib offset | -14 to –65, should settle at –15 but may take as much as an hour (don’t wait for it!). |  |  |
| calib error | 0x08, should settle to 0x00 |  |  |

j) Verify Analog Housekeeping

NOTE: Temperature housekeeping is still being calibrated  4/27/05

| Description | Expected Values | Actual Values | Comments |
| --- | --- | --- | --- |
| HV | 0-100v |  |  |
| Tof T | 25-35 C |  |  |
| Foil T | 20-25 C |  |  |
| SSD T | 20-25 C |  |  |
| +3.3v monitor | 3.3v |  |  |
| +2.5v monitor | 2.5v |  |  |
| +5v monitor | 5v |  |  |
| +6v monitor | 6v |  |  |

3.3 SIT  Quiet Mode

Commands – send the following sequence of commands:

immed 1		causes future commands to be executed upon receipt

hvenable 0		ensures that the SIT HVPS is disabled

junk 1			causes instrument to respond to all events, even illegals		eonly 1		removes TOF coincidence requirement for events

toferror 1		processes events including tof errors

limhi 100

Command Verification  - After approximately 2 minutes verify that all the above commands have been received and acted on.

| Description | Expected Value | Actual Value | Comments |
| --- | --- | --- | --- |
| Toferror | 1 |  |  |
| HV Status | 0 |  |  |
| Eonly | 1 |  |  |
| Junk | 1 |  |  |
| limhi | 256 |  |  |

3.4   Noisy Mode

Commands – send the following sequence of commands:

The following is a HAZARDOUS COMMAND

hvenable 1		enables the SIT HVPS

The following is a HAZARDOUS COMMAND – verify command carefully before sending!

hvlevel 10		sets the HVPS  output to about 300v

Command Verification  - After approximately 2 minutes, verify that all  the above commands have been received and acted on.

| Description | Expected Value | Actual Value | Comments |
| --- | --- | --- | --- |
| HV step | 10 |  |  |
| Toferror | 1 |  |  |
| HV status | 1 |  |  |
| Eonly | 1 |  |  |
| Junk | 1 |  |  |
| HV (see note*) | 300-350v |  |  |

*Note: Housekeeping for the HV Monitor should read 300-350v if converted to engineering units.  This will probably cause a Yellow Limit alarm on the GSE monitoring SIT housekeeping, but no corrective action is required.

  - Normal Mode – HV Off

Commands – send the following sequence of commands:

hvenable 0		ensures that the SIT HVPS is disabled

hvlevel 0		sets HV level to 0

junk 0			causes instrument to reject illegal events

eonly 0		removes TOF coincidence requirement for events

toferror 0		processes events including tof errors

limhi 1f4		limits hi priority events to 500/major frame

Command Verification  -  After approximately 2 minutes, log and check the following:

| Description | Expected Value | Actual Value | Comments |
| --- | --- | --- | --- |
| HV step | 00 |  |  |
| Toferror | 0 |  |  |
| HV Status | 0 |  |  |
| Eonly | 0 |  |  |
| Junk | 0 |  |  |
| LIMHI | 500 |  |  |

3.6 Pulser Test

a) Turn Pulser ON/OFF switch to ON (or turn DG535 power ON)

b) After 2-3 minutes verify data as follows:

| Description | Expected Value | Actual Value | Comments |
| --- | --- | --- | --- |
| SRT | 720 |  |  |
| STP | 6000 |  |  |
| VS | 6000 |  |  |
| SSD | 6000 - 6020 |  |  |
| VSE (Event) | ~6000 |  |  |
| SRTE (Artifical Stop) | 0-5 |  |  |
| STPE (Error) | 0-5 |  |  |

Record Typical Pulse Height Event (ApId 606-617):

tofch   range: __________________________________

ssdch  range:  __________________________________

Record Non-zero Matrix Rates (ApId 605) :  ______________

______________

______________

Beacon Rates (ApId 619):  all 0s                   _______________

c) After about 10 minutes of stable data observation, turn pulser ON/OFF switch to OFF.

3.7 End of Test

a)  Disconnect Test Pulser from SIT and store.

b)  Use SIT_Payload_File to convert data file to ascii

c)  Anonymous ftp file to uleis.umd.edu

log in as anonymous

cd ftk6/upload

ascii

put file

Fig.1a  Bench Pulser Setup

# Text boxes

# BH1

50

Minimum Delay

Green BNC (3)

0.02

2.5

50ohm Term (2)

20 dB

DG535 Setup:

Trig	Internal, 100 Hz

Delay:  A=T+0

B=A+3ns

C=A+20ns

D=C+3ns

Output:AB 50ohm,var,+1,-1

CD 50ohm, var,+1,-1

EXT

20 dB

START

STOP

BNC to Nanohex

# DG535

START

Energy

STOP

To Energy

SIT Electronics

(Front View)